# How MET/CAL Determines the 9500(B) Active Head in V6.1X

Application Information

Determination of Active Head Requirement

MET/CAL attempts to determine the minimum capability active head which is

required to execute a 9500 statement.  For example the following statement

specifies a 3.2GHz, 1Vpp, 50 Ohm Leveled-sine wave signal from the 9500:

1.001  9500         3.2GH                        1Vpp           LS    S  L

This signal requires at least a 9530 Active Head.  However at run time,

the statement will also execute if a 9560 Active Head is physically used.

In addition MET/CAL also attempts to determine the output channel that the

signal is routed to.  If no explicit output channel is specified, channel 1

is presumed.  Therefore the above statement, by itself, will generate the

the following entries in the procedure header:

CONFIGURATION:         Datron 9500 (G3)

CONFIGURATION:         Datron 9530 on channel 1

(G3) specifies that a 9500/3200 or 9500B/3200 is required.

The output channel may be specified using the M9500 FSC.

1.001  M9500                       CH2

1.002  9500         3.2GH                        1Vpp           LS    S  L

In this case, the entries in the procedure header would be:

CONFIGURATION:         Datron 9500 (C5,G3)

CONFIGURATION:         Datron 9530 on channel 2

Note the 9530 is now required on channel 2.  (C5) has also been added since

Option 5 (five channel output is required).

Now, the above process for determining the Active Head requirement works

fine as long as MET/CAL can correctly determine the output channel that

will be used at run time.  However, problems can occur when MET/CAL

procedure writers use a subprocedure to perform test which is redundant

over multiple UUT inputs.  For example a typical scenario is the following:

The UUT is a four channel oscilloscope.

The manufacturer's service manual lists the steps necessary to perform

a bandwidth test.

The manual then dictates that the test be performed on each input channel.

The procedure writer developes a channel independant subprocedure to perform

a bandwidth test.

(c) 2001 Fluke Corporation, All rights reserved             MET/CAL Procedure

=============================================================================

INSTRUMENT:            Sub Bandwidth Test

DATE:                  2001-09-12

AUTHOR:                Fluke

REVISION:              1.0

ADJUSTMENT THRESHOLD:  70%

NUMBER OF TESTS:       1

NUMBER OF LINES:       26

CONFIGURATION:         Datron 9500

CONFIGURATION:         Datron 9530 on channel 1

=============================================================================

STEP    FSC    RANGE NOMINAL        TOLERANCE     MOD1        MOD2  3  4 CON

1.001  HEAD         ANALOG BANDWIDTH CHECK - CHANNEL [M100]

#                     .

#                     .

#                     .

1.002  9500         0.600Vpp                     3MH            LS    S  L

#                     .

#                     .

#                     .

1.006  END

Note the subprocedure header only requires a 9530 on channel 1, because

no M9500 statement is included.  The M9500 statement is omitted to make

the subprocedure channel independent.

The procedure writer then writes a main procedure which calls the bandwidth

subprocedure for each UUT channel.

(c) 2001 Fluke Corporation, All rights reserved             MET/CAL Procedure

=============================================================================

INSTRUMENT:            Main

DATE:                  2001-09-12

AUTHOR:                Fluke

REVISION:              1.0

ADJUSTMENT THRESHOLD:  70%

NUMBER OF TESTS:       4

NUMBER OF LINES:       31

=============================================================================

STEP    FSC    RANGE NOMINAL        TOLERANCE     MOD1        MOD2  3  4 CON

1.001  LABEL        CH1_BANDWIDTH

1.002  M9500                       CH1

1.003  CALL         Sub Bandwidth Test

1.004  JMPL         CH2_BANDWIDTH

1.005  EVAL   dummy

2.001  LABEL        CH2_BANDWIDTH

2.002  M9500                       CH2

2.003  CALL         Sub Bandwidth Test

2.004  JMPL         CH3_BANDWIDTH

2.005  EVAL   dummy

3.001  LABEL        CH3_BANDWIDTH

3.002  M9500                       CH3

3.003  CALL         Sub Bandwidth Test

3.004  JMPL         CH4_BANDWIDTH

3.005  EVAL   dummy

4.001  LABEL        CH4_BANDWIDTH

4.002  M9500                       CH4

4.003  CALL         Sub Bandwidth Test

4.004  END

Note the main procedure header does not list a 9500 or any active heads

as required.  Now if the procedure is executed, the M9500 statements in

the main procedure will designate the output channel to be used before

each call to the subprocedure.  The 9500 run time driver will use this

information when the 9500 statement is executed to select the correct

output channel.  However when the driver attempts to designate the Active

connected to this output channel as "used", an error will be generated.

The error at this point is due to the run time driver attempting to

use a traceable asset, the active head, which was not required by the

main procedure or any subprocedure as indicated in the header.

To avoid this problem, the main or subprocedure header must contain

configuration lines for the actual output channels that will be used.

The most straight forward way to accomplish this is to include "dummy"

statements in the main procedure that force the configuration lines to

be generated.

(c) 2001 Fluke Corporation, All rights reserved             MET/CAL Procedure

=============================================================================

INSTRUMENT:            Main

DATE:                  2001-09-12

AUTHOR:                Fluke

REVISION:              1.1

ADJUSTMENT THRESHOLD:  70%

NUMBER OF TESTS:       4

NUMBER OF LINES:       51

CONFIGURATION:         Datron 9500 (C5,G3)

CONFIGURATION:         Datron 9530 on channel 1

CONFIGURATION:         Datron 9530 on channel 2

CONFIGURATION:         Datron 9530 on channel 3

CONFIGURATION:         Datron 9530 on channel 4

=============================================================================

STEP    FSC    RANGE NOMINAL        TOLERANCE     MOD1        MOD2  3  4 CON

1.001  JMPL         CH1_BANDWIDTH

1.002  M9500                       CH1

1.003  9500         3GH                          1Vpp           LS    S  L

1.004  M9500                       CH2

1.005  9500         3GH                          1Vpp           LS    S  L

1.006  M9500                       CH3

1.007  9500         3GH                          1Vpp           LS    S  L

1.008  M9500                       CH4

1.009  9500         3GH                          1Vpp           LS    S  L

1.010  LABEL        CH1_BANDWIDTH

1.011  M9500                       CH1

1.012  CALL         Sub Bandwidth Test

1.013  JMPL         CH2_BANDWIDTH

1.014  EVAL   dummy

2.001  LABEL        CH2_BANDWIDTH

2.002  M9500                       CH2

2.003  CALL         Sub Bandwidth Test

2.004  JMPL         CH3_BANDWIDTH

2.005  EVAL   dummy

3.001  LABEL        CH3_BANDWIDTH

3.002  M9500                       CH3

3.003  CALL         Sub Bandwidth Test

3.004  JMPL         CH4_BANDWIDTH

3.005  EVAL   dummy

4.001  LABEL        CH4_BANDWIDTH

4.002  M9500                       CH4

4.003  CALL         Sub Bandwidth Test

4.004  END