Measure for Measure

November, 2000

Philip Stein

Don’t Whine, Calibrate

I frequently get questions about whether some process or measurement requires calibration. Most of the time, it’s like the recent one in which a quality engineer from a printing company wrote to ask whether his densitometer (used to check colors in a print job) needed to be included in his ISO 9000 paragraph 4.11 activities.

First, I answered by quoting the standard. “The supplier shall establish and maintain documented procedures to control, calibrate, and maintain inspection, measuring, and test equipment (including test software) used by the supplier to demonstrate the conformance of product to the specified requirements.” (4.11.1)

If a measurement is not used to ‘demonstrate conformance to requirements’, it’s not covered by 4.11. In the above example, If a subsequent inspection or test procedure is used as the final arbiter of whether the color of a job is right, the densitometer in question need not be calibrated, at least according to the requirements of the standard.

But does it need to be calibrated for other reasons?

My final answer to this correspondent, and to you the reader as well, is the following simple test:

Does it matter whether the answer from this measurement is correct?

•	If it does matter, then calibration is needed.

•	If it doesn’t matter, then why are you making the measurement?

I’ve seen some cases where the value of a measurement is not important, but the trend is. When grinding lenses, for example, a dial indicator is used to enable to operator to see the depth of penetration of the lap into the blank. When progress is stopped, or nearly so, the grinding is finished. The value of the measurement on the dial is not important, only the rate of change of position of the pointer matters. In terms of the question above, it doesn’t matter whether the answer is correct, but there is still a reason to make the measurement. This instrument therefore did not need calibration and was marked ‘for reference only’.

The real goal

So with this simple test, you should be able to determine whether to calibrate. Of course the reason the engineer with the densitometer was asking was not idle curiosity. Someone, probably his boss, wanted to save money by eliminating some costly calibrations. With the test, though, you can see that it matters if the answer is wrong, so this is the wrong place to cut corners.

Cutting corners may not be a good idea here, but that doesn’t mean we can’t save money. Too often, calibration is not done economically, and costs can be reduced without compromising results. In my last column here, I said that calibration, and the choice of how often to calibrate, is dependent on economic risk. The calibration interval is an economic tradeoff in which you balance the frequency of calibration against the cost of calibration and against the potential consequences and costs of not calibrating. In some cases it might be smart to calibrate more frequently if the costs of being wrong are enormous.

Here’s a real-life situation in which we saved big bucks just by being smart and knowing something about measuring tools. One of my clients makes rubber products. They hold some dimensions of these products to tight tolerances, and therefore the correctness of readings from their hand-held micrometers is important.

To start the money-saving process, we need to know how a micrometer works and, more important, why it is likely to read incorrectly and need calibration. When a micrometer is new (of recently introduced into the system, even if it’s old), all sorts of factors such as the pitch of the screw need to be verified. Unless the tool is used hundreds of times a day to check exactly the same dimensions, screw wear doesn’t happen very fast and so doesn’t have to be checked very often. What does wear on most micrometers is the faces of the spindle and anvil. Measuring abrasive materials such as Titanium or Teak could significantly wear the tool in a matter of weeks. Measuring rubber, however, may never make a difference.

This company was calibrating micrometers monthly, regardless of how much it was used and how often. There are a few unusual circumstances in which I would have said that this wasn’t often enough, but in a rubber factory, It’s far too often. The calibration interval for micrometers used on rubber is now one year. Saved a lot of money. With some more experience, we may be able to comfortably increase the interval to two years.

Other identical micrometers in the same company are used in general machining work and wear faster. They were reduced to three- or six-month intervals at first, and will again be reduced further if the data indicate that this is wise. Note that although we have saved money, we now have to be much smarter, and treat each tool according to how it is used. This is not as easy as having the same interval for all like tools. Even better, in this company some tools stayed in the instrument cabinet for months at a time, being removed only to calibrate them again. Maintaining tools according to the amount they have been used also requires being smarter, but the savings can be immense.

Micrometers and most other tools are subject to abuse, being dropped, etc., and may not last for two years without needing repair. It would be wrong to set an interval for two years if tools broke or bent more frequently than that. This leads us to a further consideration when deciding on not just when to calibrate, but what to do each time.

If you really do care about the results of a measurement, you might want quite a bit of confidence in the stability and reliability of the tool you use. This confidence can be had at the cost of frequent calibrations, but it can be had more cheaply as well.

Continuing our exposition of the micrometer: we determined that under normal use, the device wears slowly and needs calibration infrequently. It may break or become damaged, though, and that happens catastrophically and at any time. How do we guard against this while keeping costs to a minimum?

The answer is ‘checks’. Let’s use our understanding of how a micrometer works to develop an inexpensive procedure that will verify its operation without a full calibration. The simplest check is to run the spindle out to the end of its travel, then back to zero. If the motion feels smooth, without catching, binding, or rattling loosely, the check has passed. A second check is to close the jaws and see if the instrument reads zero (close the jaws of a larger tool onto a gage block). A digital mike can be forced to read zero at this point, so you have to look at the scale on the spindle. These two checks can be made weekly, or even before every shift. They only take a minute or less. If they pass muster, it’s likely that everything is okay. If not, send them for repair and a full recalibration.

Finally, the last act of the check is to document that it has been done. Since no numerical data are taken, a control chart would probably be overkill. Marking the date on a calendar, or on a wall sheet that covers many instruments, will produce a permanent quality record. Measuring the reliability of the tool in this way is possible, but it’s not the point. If your check finds a problem, the tool will be taken out of service, so it’s unlikely that a broken tool will ever be used. The point here is to keep it simple and cheap, and to just monitor to see that nothing catastrophic has happened. In this way you have decreased your cost, decreased your calibration frequency and increased your confidence at the same time.